FLUKE 45

Dual Display Logger R07

User Installation and Operating Manual

| Application release | R07 |
| --- | --- |
| Application version | 1.4.0 |
| Target operating system | Windows 11, 64-bit |
| Target Python | Python 3.13 |
| Document date | 31 July 2026 |

Purpose: This manual explains how to install, connect, operate, and troubleshoot Fluke 45 Logger R07. It covers live display monitoring and timestamped CSV data logging through the Fluke 45 rear RS-232 interface.

For laboratory, R&D, service, and controlled measurement use

# Document Control

| Item | Description |
| --- | --- |
| Application | Fluke 45 Dual Display Logger R07 |
| Software version | 1.4.0 |
| GUI framework | Python standard Tkinter and ttk |
| External runtime dependency | pySerial 3.5 |
| Communication interface | Fluke 45 rear RS-232 serial port |
| Data format | Comma-separated values (CSV) |
| Timestamp source | Windows PC local clock, millisecond resolution |

# Contents

1. Application Overview

2. System and Hardware Requirements

3. Installation

4. Fluke 45 Serial Setup

5. Starting the Application

6. User Interface Guide

7. Connecting and Identifying the Meter

8. Live Display Operation

9. CSV Logging Operation

10. CSV File Structure

11. Recommended Operating Practices

12. Diagnostics and Troubleshooting

13. Configuration, Logs, and File Locations

14. Shutdown, Reinstallation, and Removal

15. Acceptance Checklist

Important R07 behavior: Continuous live display and CSV logging are intentionally mutually exclusive. Stop live display mode before starting CSV logging. During CSV logging, each logged sample also updates the primary and secondary GUI displays.

# 1. Application Overview

Fluke 45 Logger R07 is a Windows desktop application designed to communicate with a Fluke 45 dual-display digital multimeter through its RS-232 port. The application identifies the connected meter, reads the primary and secondary displays, provides continuous live monitoring, and records measurements to CSV files.

## 1.1 Main Functions

Automatic Windows COM-port discovery and refresh.

Fluke 45 identification using the instrument identification query.

Primary and secondary display acquisition.

One-shot acquisition using Acquire Single Reading.

Continuous live display with a user-defined update interval.

Timed or continuous CSV logging.

Independent live-display and CSV sample intervals.

CSV duration entered in minutes, with 0 meaning continuous logging.

Local PC timestamps with millisecond resolution.

Slow, Medium, and Fast meter-rate selection for CSV sessions.

Background serial acquisition so the GUI remains responsive.

Application event log plus persistent rotating diagnostic logs.

## 1.2 Architecture and Data Integrity

The serial port is owned by one background worker thread. The GUI does not directly compete for the COM port. This design prevents overlapping serial commands and reduces the risk of prompt or response misalignment. CSV files are created exclusively so an existing file is not silently overwritten, and every successful sample is flushed to disk immediately.

# 2. System and Hardware Requirements

| Requirement | R07 Requirement |
| --- | --- |
| Operating system | Windows 11, 64-bit |
| Python | Python 3.13 installed with the Python Launcher |
| Tkinter | Python Tcl/Tk and IDLE component installed |
| Serial library | pySerial 3.5, installed by the R07 installer |
| Meter | Fluke 45 digital multimeter |
| Cable | RS-232 null-modem or crossover serial cable |
| PC serial interface | Native RS-232 or USB-to-RS-232 adapter |
| Not suitable | USB-to-TTL serial adapters |

## 2.1 Hardware Connection

Fluke 45 rear DB-9 RS-232 port ¦ RS-232 null-modem / crossover cable ¦ USB-to-RS-232 adapter, if required ¦ Windows 11 computer

Electrical interface: Use a true RS-232 interface. A USB-to-TTL adapter uses different electrical levels and is not appropriate for the Fluke 45 RS-232 port.

# 3. Installation

## 3.1 Extract the Package

1.  Extract the R07 ZIP package into a normal Windows folder. Do not run the application directly from inside the ZIP file.

2.  A typical folder can be:

C:\AI\Fluke45 SerialCom\Fluke45_Logger_R07

3.  If upgrading from an earlier release, do not copy the older .venv folder into R07.

## 3.2 Run the Installer

1.  Double-click install_fluke45.bat.

2.  The installer checks Python 3.13 and the Tcl/Tk runtime.

3.  It creates a private .venv inside the application folder.

4.  It installs pySerial 3.5 into that private environment.

5.  It verifies application imports and constructs the GUI in hidden mode.

6.  It runs the R07 automated test suite.

Successful installation: The final message should state that installation and self-tests completed successfully and that the application can be started using run_fluke45.bat.

## 3.3 Verify Python Tkinter if Installation Fails

If the installer reports that Tcl/Tk cannot be found or a Tk window cannot be created, test the base Python installation from Command Prompt:

py -3.13 -m tkinter

A small Tk demonstration window should open. If it does not, use the Python 3.13 Windows installer Modify or Repair option and ensure the tcl/tk and IDLE feature is enabled.

## 3.4 Start the Application

1.  Double-click run_fluke45.bat.

2.  The Fluke 45 Dual Display Logger window opens.

3.  The application remembers valid user settings such as the last COM port, baud rate, output folder, intervals, meter rate, and appearance.

# 4. Fluke 45 Serial Setup

The application and the Fluke 45 must use matching serial settings. A recommended starting point is:

| Parameter | Recommended Setting |
| --- | --- |
| Interface | RS-232 |
| Baud rate | 9600 |
| Parity | None |
| Stop bits | 1 |
| Echo | OFF |
| Print-only rate | 0 |

R07 supports Fluke 45 baud rates of 300, 600, 1200, 2400, 4800, and 9600 baud. When parity is None, the application uses 8 data bits. When parity is Even or Odd, the application uses the meter-compatible 7-bit framing with one stop bit.

## 4.1 Finding the Windows COM Port

1.  Connect the USB-to-RS-232 adapter to the computer.

2.  Open Windows Device Manager.

3.  Expand Ports (COM & LPT).

4.  Record the COM number assigned to the adapter.

5.  In R07, use Refresh if the desired port is not already shown in the COM port list.

# 5. Starting the Application

At startup, the application scans the available Windows serial ports. The top of the window shows connection controls on the left and the live primary and secondary meter displays on the right.

Recommended first-use sequence: Set the Fluke 45 serial interface, connect the cable, start R07, select the correct COM port, confirm 9600 baud and matching parity, then use Connect and Identify before starting live display or CSV logging.

# 6. User Interface Guide

| GUI Section | Purpose | Main Controls |
| --- | --- | --- |
| 1. Serial Connection | Select and establish the RS-232 connection. | COM port, Baud rate, Parity, Refresh, Connect and Identify, Disconnect, instrument identity. |
| 2. Live Meter Display | View primary and secondary Fluke 45 readings. | Live interval (s), Acquire Single Reading, START LIVE DISPLAY, STOP LIVE DISPLAY. |
| 3. CSV Logging Session | Configure and run a measurement recording session. | Test name, Output folder, CSV interval (s), Duration (min), Meter rate, START CSV LOGGING, STOP LOGGING. |
| 4. Application Event Log | Display connection, acquisition, logging, and error events. | Time-stamped status messages and operating history. |

## 6.1 Connection Status

The status indicator at the top-right of the application shows whether the software is connected to the meter. After successful identification, the instrument panel displays the manufacturer, model, serial number, and firmware information returned by the Fluke 45.

## 6.2 Display Fields

The PRIMARY DISPLAY and SECONDARY DISPLAY areas show the parsed numeric values and unit mnemonics returned by the meter. If the secondary display is not active on the Fluke 45, R07 shows the appropriate inactive state rather than inventing a second measurement.

# 7. Connecting and Identifying the Meter

1.  Power on the Fluke 45.

2.  Confirm the rear RS-232 cable is connected.

3.  Select the correct COM port.

4.  Select the baud rate configured on the meter.

5.  Select the matching parity.

6.  Click Connect and Identify.

R07 opens the COM port, synchronizes with the Fluke 45 command prompt, queries the instrument identity, validates that the connected device is a Fluke 45, and configures measurement responses with units.

If identification fails: Check the COM port, cable type, baud rate, parity, Fluke 45 RS-232 mode, and whether another Windows program currently has the COM port open.

# 8. Live Display Operation

## 8.1 Single Reading

1.  Connect and identify the meter.

2.  Set the desired measurement function on the Fluke 45.

3.  Click Acquire Single Reading.

The application requests the current display value or values and updates the GUI once. No CSV row is written.

## 8.2 Continuous Live Display

1.  Connect and identify the meter.

2.  Enter the required Live interval (s). A practical starting value is 1 second.

3.  Click START LIVE DISPLAY.

4.  The primary and secondary display panels update continuously at the requested interval, subject to meter and serial response time.

5.  Click STOP LIVE DISPLAY when continuous monitoring is no longer needed.

The accepted live interval range is 0.05 seconds to 86,400 seconds. Selecting a very short interval does not force the Fluke 45 to produce measurements faster than its own measurement rate and serial transaction time allow.

Automatic error protection: Five consecutive protocol acquisition errors stop the active live display mode. A serial-link failure also stops acquisition and releases the connection path.

# 9. CSV Logging Operation

## 9.1 Configure the Session

1.  Stop continuous live display if it is running. Live display and CSV logging cannot run at the same time.

2.  Enter a Test name. R07 uses this as part of the CSV filename and sanitizes characters that are unsafe in Windows filenames.

3.  Use Browse to select the Output folder.

4.  Enter CSV interval (s). This controls the requested time between saved samples.

5.  Enter Duration (min). Use 0 for continuous logging until STOP LOGGING is pressed.

6.  Select the Meter rate: Slow, Medium, or Fast.

7.  Click START CSV LOGGING.

During a CSV session, each successful logged sample also refreshes the primary and secondary GUI displays. The Samples counter, Elapsed time, Errors counter, and active CSV filename are updated in the session panel.

## 9.2 Duration Examples

| Duration Entry | Result |
| --- | --- |
| 0 | Continuous logging until STOP LOGGING is pressed |
| 1 | 1 minute |
| 15 | 15 minutes |
| 60 | 60 minutes |
| 120 | 2 hours |

## 9.3 Stopping a Session

Press STOP LOGGING to close the active CSV session. The application flushes and closes the file. You can then use Open Output Folder to open the configured Windows folder.

# 10. CSV File Structure

R07 writes local PC time only. It does not create a UTC timestamp column.

| Column | Description |
| --- | --- |
| sample_index | One-based sequence number for logged samples. |
| local_timestamp | Windows PC local time with time-zone offset and millisecond resolution. |
| elapsed_s | Elapsed session time in seconds. |
| primary_value | Primary display numeric value. |
| primary_unit | Primary display unit returned by the meter. |
| secondary_value | Secondary display numeric value when active. |
| secondary_unit | Secondary display unit when active. |
| raw_response | Original parsed Fluke 45 response text. |

## 10.1 Example

The CSV excerpt below may wrap visually to fit the page width. The actual CSV file stores each sample as one row.

sample_index,local_timestamp,elapsed_s,primary_value,primary_unit,secondary_value,secondary_unit,raw_response 1,2026-07-30T20:15:00.125-06:00,0.000,0.567,OHMS,,,"+5.6700E-1 OHMS" 2,2026-07-30T20:15:02.126-06:00,2.001,0.568,OHMS,,,"+5.6800E-1 OHMS"

Files are written as UTF-8 with a byte-order mark for good spreadsheet compatibility. The writer also prevents accidental file overwrite and guards formula-like text fields when the CSV is opened in spreadsheet software.

# 11. Recommended Operating Practices

Start with 9600 baud and parity None when the meter is configured that way.

Use a 1-second live interval for general monitoring, then shorten it only if needed.

Use a CSV interval that is realistic for the selected Fluke 45 measurement rate.

Confirm the Windows PC clock before tests where timestamp traceability matters.

Use descriptive test names such as Supply_24V_Thermal or Resistance_Baseline.

Write CSV files to a local folder with reliable write access.

Stop logging before disconnecting the serial cable or powering off the meter.

Use the Application Event Log to confirm connection, acquisition start, and stop events.

## 11.1 Live Display Versus Logging

| Operating Mode | Writes CSV | Updates GUI | When to Use |
| --- | --- | --- | --- |
| Acquire Single Reading | No | Once | Manual spot checks |
| START LIVE DISPLAY | No | Continuously | Real-time observation |
| START CSV LOGGING | Yes | Each logged sample | Recorded test sessions |

# 12. Diagnostics and Troubleshooting

| Symptom | Likely Cause | Corrective Action |
| --- | --- | --- |
| COM port not listed | Adapter not detected or port list is stale | Check Device Manager, reconnect the adapter, then click Refresh. |
| Connect and Identify fails | Wrong COM port, baud, parity, cable, or interface mode | Match the Fluke 45 serial settings and verify RS-232 cabling. |
| Port access error | Another application is using the COM port | Close terminal, vendor, or logging software using that port. |
| Live display stops | Repeated protocol errors or serial interruption | Review the event log, verify cabling, then reconnect. |
| Secondary display is inactive | Fluke 45 secondary display is not enabled | Enable the desired secondary measurement on the meter. |
| CSV will not start | Live display is still active or output path is invalid | Stop live display and verify the selected output folder. |
| CSV interval seems slower than requested | Meter conversion or serial transaction time is limiting | Use a realistic interval or select a faster meter rate. |
| Tkinter installation error | Python Tcl/Tk component is unavailable | Run py -3.13 -m tkinter and repair Python Tcl/Tk if required. |

## 12.1 Diagnostic Batch File

R07 includes:

diagnose_fluke45.bat

Use it when investigating Python, Tcl/Tk, dependency, or application startup problems.

# 13. Configuration, Logs, and File Locations

| Item | Location / Behavior |
| --- | --- |
| Default CSV folder | %USERPROFILE%\Documents\Fluke45_Logs |
| Application log | %LOCALAPPDATA%\Fluke45Logger\logs\fluke45_application.log |
| Launcher error log | <R07 application folder>\launcher_error.log |
| Persistent settings | %LOCALAPPDATA%\Fluke45Logger\config.json |
| Private Python environment | <R07 application folder>\.venv |

The application event log shown in the GUI is bounded so long sessions do not grow the display indefinitely. Persistent diagnostic logging uses the Windows local application-data area.

# 14. Shutdown, Reinstallation, and Removal

## 14.1 Normal Shutdown

1.  Stop live display if active.

2.  Stop CSV logging if active.

3.  Disconnect the Fluke 45.

4.  Close the application window.

## 14.2 Reinstallation

For a clean R07 reinstall, close the application, delete the R07 .venv folder, then run install_fluke45.bat again. The application settings stored under LocalAppData are separate from the private environment.

## 14.3 Removal

Because R07 is folder-based, removal is straightforward. Close the application, delete the extracted R07 application folder, and optionally delete the Fluke45Logger folder under LocalAppData if you also want to remove saved settings and diagnostic logs.

# 15. Acceptance Checklist

| # | Verification | Pass |
| --- | --- | --- |
| 1 | R07 package extracted to a normal Windows folder | ☐ |
| 2 | install_fluke45.bat completes successfully | ☐ |
| 3 | run_fluke45.bat opens the GUI | ☐ |
| 4 | Correct COM port appears after Refresh | ☐ |
| 5 | Connect and Identify displays Fluke 45 identity | ☐ |
| 6 | Acquire Single Reading updates the primary display | ☐ |
| 7 | Secondary display updates when enabled on the meter | ☐ |
| 8 | START LIVE DISPLAY updates continuously | ☐ |
| 9 | STOP LIVE DISPLAY stops acquisition cleanly | ☐ |
| 10 | CSV test session starts and creates a new file | ☐ |
| 11 | CSV timestamps match the Windows PC local clock | ☐ |
| 12 | Duration in minutes stops the session as expected | ☐ |
| 13 | STOP LOGGING closes the file cleanly | ☐ |
| 14 | CSV opens correctly in spreadsheet software | ☐ |

Final equipment verification: Software validation for R07 included Python compilation, hidden GUI construction, parser and serial-transaction simulation, continuous monitor tests, CSV schema tests, and 38 automated tests. Final acceptance should still confirm sustained operation using the actual Fluke 45, serial cable, USB-to-RS-232 adapter, and selected sampling interval.

# Appendix A - Quick Reference

| Task | Action |
| --- | --- |
| Install | Run install_fluke45.bat once after extraction. |
| Launch | Run run_fluke45.bat. |
| Connect | Select COM, baud, parity, then Connect and Identify. |
| One reading | Click Acquire Single Reading. |
| Live monitor | Set Live interval (s), then START LIVE DISPLAY. |
| Stop live monitor | Click STOP LIVE DISPLAY. |
| Log to CSV | Configure session, then START CSV LOGGING. |
| Continuous CSV | Set Duration (min) to 0. |
| Stop logging | Click STOP LOGGING. |
| Find results | Click Open Output Folder. |
| Diagnostics | Run diagnose_fluke45.bat and review log files. |

# Appendix B - Recommended First Test

1.  Set the Fluke 45 for RS-232, 9600 baud, parity None, echo OFF.

2.  Connect the RS-232 link and start R07.

3.  Connect and identify the meter.

4.  Acquire one manual reading.

5.  Set Live interval to 1 second and run live display for 30 seconds.

6.  Stop live display.

7.  Set Test name to Initial_Check.

8.  Set CSV interval to 2 seconds and Duration to 1 minute.

9.  Start CSV logging and allow the session to complete.

10.  Open the output folder and verify the CSV values and PC timestamps.